Stability-indicating HPLC-DAD method for the determination of ipragliflozine

Fathy M. Salama, Khalid A. M.  Attia, Ahmed A. Abouserie, Ragab A. Mabrouk, Ahmed M Abdelzaher *

Pharmaceutical Analytical Chemistry Department, Faculty of Pharmacy, Al-Azhar University, 11751 Nasr City, Cairo, Egypt.

*Corresponding author. Mobile: +201002845286,

E-mail address:   Ahmedabdelzaher86@yahoo.com,  Ahmedabdelzaher86@azhar.edu.eg

## Abstract

## A novel stability-indicative of RP-HPLC method was developed and validated for the quantitative estimation of ipragliflozin in bulk drugs and pharmaceutical dosage form in the presence of its forced degradation products. The drug was subjected to variable stress conditions including hydrolysis, oxidation, thermal and photolysis. The drug was found to be labile to acidic hydrolysis, basic hydrolysis, oxidation but stable in thermal, photolysis and neutral hydrolytic conditions. Successful chromatographic separation of ipragliflozin was achieved on HAMILTON column prp-x100 C18 column at a flow rate of 1 mL/min using mobile phase composition of methanol : acetonitrile : Water in ratio of (50: 25 : 25, by volume). The eluents were monitored by the diode array detector and peak area values were measured at 230 nm. The validity of the method was assessed by evaluating accuracy, precision, speciﬁcity and robustness. The linear regression analysis data for the calibration curve shows a good relationship in the range of  10 – 100 µg/mL. The developed method was successfully applied for the estimation of  ipragliflozin in its commercial dosage form and could be used for the routine analysis of the studied drug in quality control laboratories.

## .Keywords

Ipragliflozine; HPLC-DAD; Stability-indicating assay; Degradation product.

## 1. Introduction

Ipragliflozine,{ (1S)-1,5-anhydro-1-C-{3-[(1-benzothiophen-2-yl)methyl]-4-fluorophenyl}-D-glucitol}, ( Suglat®, Fig. 1) is a highly potentand selective sodium-dependent glucose co-transporter-2 (SGLT2)inhibitor, a novel class of hypoglycemic agents [1–4]. It is white powder practically insoluble in water and propylene glycol; very slightly soluble in ethanol; slightly soluble in acetone and soluble in methanol [5].  SGLT2 inhibitors suppress glucose re-absorption in the proximal tubules of the kidney and promote the urinary excretion of glucose; there-fore, SGLT2 inhibitors may reduce blood glucose levels and bodyweight [4]. Tahara et al. previously reported that ipragliflozin improved diabetes/obesity-associated metabolic abnormalities in type 2 diabetic mice, which suggests that ipragliflozin may be valuable in the treatment of type 2 diabetes with metabolic syndrome [5]. Additional advantages of SGLT2 inhibitors aredecreases in blood pressure, the prevention of vascular complications, and recovery of insulin sensitivity and  β-cell function[6].

## The analytical methods, including high performance liquid chromatography (HPLC) with an ultraviolet detector [5,7] and LC–MS/MS [8,9] have been used in quantitative analysis of ipragliflozin in different biological matrices.

## In this paper, we described validation of a method for accurate quantification of an ipragliflozin in bulk drugs and pharmaceutical dosage form along with validation as per ICH [10-12]. The plan of present study was to establish the inherent stability of ipragliflozin through stress studies under variety of conditions and develop a stability –indicating RP-HPLC method.

Fig. 1: Structure formula of ipragliflozin.

## 2. Experimental

### 2.1. Instruments

HPLC, constaMetric® 4100 LDC Analytical pump (Milton Roy, USA), equipped with SpectraSYSTEM UV3000 Diode-array UV-Visible detector and SpectraSYSTEM AS3000 auto sampler. The chromatographic analysis was carried out using (ChromQuest 4.2.34, version 3.1.6) data analysis program.

### 2.2 Samples

- ipragliflozine (99.8%) was kindly supplied by Al Andalous for Pharmaceutical Industries, Obour city, Egypt.

### 2.3. Chemicals and solvents

- Acetonitrile and methanol, HPLC grade (Sigma-Aldrich, Germany)

- Sodium hydroxide, hydrochloric acid and hydrogen peroxide, analytical grade (El-Nasr Pharmaceutical Chemicals Co., Egypt).

### 2.4. General procedure

2.4.1 Chromatographic conditions

A mobile phase system of methanol, acetonitrile and Water in ratio of (50: 25 : 25, by volume) was used at a flow rate of 1 mL/min. The injection volume was 20 µL. The column used was HAMILTON column prp-x100 C18 column (15 cm X 4.6 mm, 5μm particle size). The eluents were monitored by the diode array detector (DAD) from 210 to 300 nm and peak area values were measured at 230 nm. All determinations were performed at ambient temperature.

2.4.2. Preparation of standard solutions and construction of calibration graph

Stock solutions (100 µg/mL) of ipragliflozin  was prepared in methanol. The working standard solutions were prepared by dilution of aliquots of the stock solutions with methanol to reach the concentration ranges of 10–100 µg/mL. The stock and working solutions were kept away from light to prevent possible photodegradation. Triplicate injections (20 µL) were made for each concentration and analyzed under the previously described chromatographic conditions. Peak areas at 230 nm were plotted against the corresponding concentrations to construct the calibration graph and the linear regression equations was derived.

2.4.3. Assay of tablets dosage form

The content of seven Suglat® Tablets (50 mg of ipragliflozin  per tablets) were mixed and weighed, after removeing the film coated by scratching and washing with methanol, then a quantity equivalent to 10 mg of ipragliflozin  was accurately transferred to 100-mL volumetric flask and the volume was made up to 75 mL with methanol. The solution was shaken vigorously for 15 min then sonicated for 30 min and completed to 100 mL with the same solvent to obtain a solution labeled to contain 100 μg/mL of ipragliflozin . The solution was filtered with 0.45 µm membrane filter then necessary dilutions of the filtrate were made with methanol to obtain different concentrations of ipragliflozin . The chromatographic conditions were applied for the prepared solutions and ipragliflozin  content of Tablets was calculated using the corresponding regression equation.

2.4.4. Preparation of forced degradation solutions

Stock standard solution (100 μg/mL) was used during forced degradation studies, and each degraded sample with a concentration of 80 μg/mL was prepared in methanol. The prepared solutions were kept away from light to prevent possible photodegradation. Then, the procedure described under chromatographic conditions was followed. From the peak area of ipragliflozin  in each chromatographed sample, the % degradation was then calculated. In each sample two mL of ipragliflozin  stock standard solution was added to 10-ml volumetric flask.

- Acidic and basic hydrolysis were carried out using 2 mL of 0.1 N HCl, and 0.1 N NaOH for 12 and 24 hours at room temperature, respectively. Each sample was neutralized with alkali or acid before dilution.

- Neutral hydrolysis was done using 2 ml deionized water for 5 hours at 80 °C.

- Oxidative degradation was carried out using 1 mL 10 % H2O2 for 12 hours at room temperature.

- Photolytic study was done by exposing solutions of the drug to sunlight for 24 hours.

- Thermal study was carried out for 12 hours at 80°C.

After the previous treatments, all forced degradation solutions were filtered with 0.45 µm membrane filters prior to injection to the column.

### 3. Results and discussion

To date, no stability-indicating method was found in the literature for the determination of ipragliflozin. In this manuscript, the goal was to develop a new stability-indicating HPLC-DAD method for the determination of ipragliflozin  in the presence of its possible degradation products.

### 3.1. Method development

A liquid chromatographic method coupled with diode array detection was developed to provide a suitable procedure for the stability study of ipragliflozin . The most important goal in liquid chromatographic method optimization is the achievement of sufficient resolution of the analytes with acceptable peak symmetry in a reasonable analysis time. The developed method was carefully designed and optimized to separate the cited compound from its forced degradation products.

To achieve this goal, three parameters should be optimized carefully. The first parameter is the reasonable choice of mobile phase. Several experiments were carried out using methanol, water and acetonitrile solutions in different ratios. It is noteworthy that ipragliflozin  is relatively more hydrophobic with molecular weight of 404.45 which is considered a macromolecule. Using combinations of water and methanol increase the retention time of ipragliflozin . It has been noted that using water in small ratios with acetonitrile lead to increase in the retention time with good resolution from its degradation products. By using mobile phase composition of methanol : acetonitrile : Water in ratio of (50: 25 : 25, by volume) a good separation was achieved with acceptable peak symmetry in a reasonable analysis time. The second parameter is the stationary phase selection. Diﬀerent stationary phases were also tried such as BDS Hypersil C18 column, Discovery® HS C18 column and HAMILTON column prp-x100 C18 column, these stationary phases gave near the results. The third parameter is the choice of the optimum wavelength at which the determination of ipragliflozin was done. The diode array detector allows the determination of ipragliflozin at different wavelengths simultaneously. Hence several wavelengths were tried but 230 nm which is the λmax of ipragliflozin  was chosen as this wavelength gave the best results regarding accuracy and sensitivity. The optimized chromatographic conditions showed a symmetric peak with minimum tailing and a reasonable analysis time. Figure 2 shows a typical chromatogram for ipragliflozin  with retention time of 3.5 ± 0.05 min. For more confirmation of the separation process the UV absorption curve of ipragliflozin  have been recorded and it shows typical matching with the curve extracted from the DAD as shown in (Fig. 3).

Fig. 2:  HPLC chromatograms of 80 μg/mL ipragliflozin  at 230 nm

Fig. 3:  The zero order UV absorbtion curve of 12 μg/mL of ipragliflozin.

### 3.2: Forced degradation studies

Forced degradation studies were carried out on ipragliflozin  to evaluate its degradation behaviour by monitoring the chromatograms of ipragliflozin  degradation products using the developed method. Due to the recent launch of this class of a novel class of hypoglycemic agents, different stress conditions were applied to evaluate the stability of ipragliflozin . It should be noted that all the recorded spectra during peak elution have been monitored using the diode array detector to detect any degradation present in the samples over the entire scanning area and to check the purity of the eluted peaks.

Ipragliflozin  was found to be sensitive to hydrolysis by 0.1 N HCL for 24 hour at room temperature with appearing of one degradation product at 1.82 min (Fig. 4). Basic hydrolysis with 0.1 N NaOH for 24 hours at room temperature with appearing of one degradation product at 1.51 min (Fig. 5). Neutral hydrolysis was tried using deionized water for 12 and 24 hours at room temperature and at 80 °C and no degradation products were found.

Fig. 4:  HPLC chromatogram of ipragliflozin and its acidic degradation product using 0.1 N HCl for 24 hour at room temperature.

Fig. 5:  HPLC chromatogram of ipragliflozin and its basic degradation product using 0.1 N NaOH for 24 hours at room temperature.

Oxidative degradation was tested using 10 % H2O2 for 12 hours. Samples were evaporated to expel the remaining H2O2. No degradation products were formed after 6 hour. However, after 12 hours one degradation product at 1.72 min was formed with complete resolving of ipragliflozin  peak (Fig. 6). Great attention has been considered for the photo degradation study of ipragliflozin . The sample of ipragliflozin  was exposed to sunlight for 6 , 12 and 24 hours. No degradation products were formed and no reduction in peak area of ipragliflozin  was observed after 6 , 12 and 24 hours. Ipragliflozin  was found to be thermally stable, as no additional peaks were observed when the drug was subjected to 80°C for 12 and 24 hours. The results of ipragliflozin  stability studies are given in (Table 1).

Fig. 6:  HPLC chromatogram of ipragliflozin and its oxidative degradation product using 10 % H2O2 for 12 hours at room temperature.

Table 1: Summary of forced degradation studies

| Stress conditions | Number of degradates (tR) | % Degradation |
| --- | --- | --- |
| 0.1 N HCL at room temperature for 24 hours | 1 (1.82) | 12.21 |
| 0.1 N NaOH at room temperature for 24 hours | 1 (1.51) | 9.43 |
| Neutral hydrolytic degradation at 80 °C for 24 hours | 0 | 0 |
| 10 % H2O2 at room temperature for 12 hours | 1 (1.72) | 24.5 |
| Sunlight for 24 hours | 0 | 0 |
| Thermal at 80 °C for 24 hours | 0 | 0 |

3.3. System suitability

System suitability was ‘checked by calculating different parameters including number of theoretical plates, Height equivalent to a theoretical plate, retention factor, resolution and tailing factor. The obtained values were found to be in the acceptable ranges when compared to USP reference values [13] as shown in (Table 2).

Table 2: System suitability results for the determination of ipragliflozin by the proposed HPLC method:

| Parameters | Ipragliflozin | Reference value [14] |
| --- | --- | --- |
| Retention time (tR) | 3.5 ± 0.05 | --- |
| Retention factor (K´) | 1.21 | 1-10 |
| Theoretical Plates (N) | 2338 | 2<000 |
| Height equivalent to theoretical plates (H in cm) | 0.011 | The smaller the value the higher the column efficiency |
| Tailing factor (T) | 1.201 | > 2 |
| Injection precision (%RSD) | 0.583 | > 1 |
| Resolution (Rs) | 0.1 N HCL (24 hour) | 1.82 | 1 < |
|  | 0.1 N NaOH (24 hours) | 1.51 |  |
|  | 10 % H2O2 (12 hours) | 1.72 |  |

### 3.4. Method validation

The proposed HPLC-DAD method was validated according to the USP on validation of analytical procedures [13]. The linearity of the developed method was determined and the linear regression data for the calibration curve showed good linearity (r2 = 0.9999) over the concentration range of 10–100 μg/mL. The accuracy of the method was calculated by measuring the mean percent recovery of three determination for three concentration of pure ipragliflozin  and it was found to be 100.12. The precision of the developed method was checked by measuring % relative standard deviation (%RSD) of three concentrations for ipragliflozin  repeated three times in the same day (repeatability) and in three days (intermediate precision). The obtained values of %RSD were < 2%, confirming the high precision of the developed method.

The specificity of the developed method was assessed by resolving ipragliflozin  from its possible degradation products. The results revealed that the proposed method was able to completely discriminate ipragliflozin  from all of its degradation products, confirming the specificity of the method. Moreover, the peak purity was checked using DAD and the purity of ipragliflozin  was found to be more than 0.992 indicating that no additional peaks were co-eluted with the main compound.

Robustness of the developed method was studied by making small changes in flow rate (±0.1 mL/min) and the working wavelengths (± 2 nm). These variations did not have any significant effects on the peak area or retention time of ipragliflozin  and %RSD of the responses were < 2% confirms the robustness of the method. The summary of these parameters and other validation and regression parameters are shown in (Table 3).

Table 3. Validation and regression parameters of ipragliflozin by the proposed HPLC method:

| Parameters | Proposed HPLC method |
| --- | --- |
| Wavelength (nm) | 230 |
| Linearity range (μg/mL) | 10 ─ 100 |
| - Regression Equation - Slope (b) - Intercept (a) | y a = b x b + a 29636 1394.9 |
| Coefficient of determination (r2) | 0.9999 |
| Accuracy | 100.61 |
| Precision (%RSD) | Repeatability | 1.063 |
|  | Intermediate precision | 1.180 |
| Robustness (%RSD) | 0.749c |
|  | 1.449d |
| LOD (µg/mL) | 0.3508 |
| LOQ (µg/mL) | 1.1694 |

a Peak area of ipragliflozin.

b Concentration of ipragliflozin.

c Robustness was checked by varying the flow rate (±0.1 mL/min).

d Robustness was checked by varying the wavelength of detection (±2 nm).

### 3.5. Application of the developed method

The developed method was applied for the assay of the cited drug in its commercial pharmaceutical formulation (Suglat® Tablets). No interfering peaks were observed from any of the excipients that may be present in the assayed preparation. The diode array detection enables peak purity verification where no signs of co-elution from any of the excipients were detected.

The standard addition technique was applied, where pure ipragliflozin  was added to already analyzed pharmaceutical preparation. Good recovery of the pure standard was obtained confirming that the proposed method could be adopted for the determination ipragliflozin  in pharmaceutical dosage form without any possible interference from excipient or additives. The obtained results are shown in (Table 4).

Table 4. Standard addition technique of the proposed HPLC method.

| Pharmaceutical taken ( µg/mL) | Pharmaceutical found ( µg/mL) | Pure added (μg/mL) | Pure found (µg/mL) | %Recovery |
| --- | --- | --- | --- | --- |
| 10 | 10.05 | 10 | 9.984 | 99.84 |
|  |  | 40 | 40.125 | 100.31 |
|  |  | 80 | 80.170 | 100.21 |
| Mean |  |  | 100.12 |
| %RSD |  |  | 0.2742 |

The obtained results of the developed method compared with the manufacturing method for determination of ipragliflozin  by HPLC  using Agilent TC (C18) column by isocratic elution of mobile phase with flow rate of 1.5 ml/min. The mobile phase composition was methanol -acetonitrile - phosphate buffer pH=3 (60:20:20, by volume). The eluate was monitored at 230 nm [14] are shown in (Table 5). The calculated t and F values are less than the tabulated ones indicating that there are no significant differences between the proposed and the reported method within a probability of 95% between the two methods.

Table 5. Determination of ipragliflozin in Suglat® Tablets by the proposed HPLC method and the reported method.

| Parameters | Proposed method | Reported method a (14) |
| --- | --- | --- |
| Mean | 100.10 | 99.58 |
| %RSD | 0.391 | 0.299 |
| Number of measurements | 5 | 5 |
| Student's t-test | 0.455 (2.306)b | —— |
| F-value | 1.224 (6.388)b | —— |

a The manufacturing method for determination of ipragliflozin  by HPLC  using Agilent TC (C18) column by isocratic elution of mobile phase with flow rate of 1.5 ml/min. The mobile phase composition was methanol -acetonitrile - phosphate buffer pH=3 (60:20:20, by volume). flow rate (1mL/min) and UV-detection at 230 nm.

b The values in parenthesis are the tabulated values of “t “and “F” at (P = 0.05)

### 4. Conclusion

This study described a simple, selective and reliable HPLC–DAD method for the estimation of ipragliflozin . To date, no comprehensive stability indicating assay method was developed for this drug by any analytical methodology. This work focused on the stability of the studied compound and its degradation behavior in different stress conditions. Significant advantages over the previously published literature were the simple mobile phase and the minimum sample preparation required for the assay of ipragliflozin . The method was validated according to the USP recommendation and can be used for the routine analysis and for checking quality during stability studies of pharmaceutical preparations containing the cited compound.

5. Conflict of interest

The authors declare no conflict of interest.

### 6. References

1. T. Kadokura, W. Zhang, W. Krauwinkel, S. Leeflang, J. Keirns, Y. Taniuchi, I.Nakajo, R. Smulders, Clinical Pharmacokinetics and Pharmacodynamics of the Novel SGLT2 Inhibitor Ipragliflozin. Clin. Pharmacokinet. 53 (2014) 975–988.

2. M. Imamura, K. Nakanishi, T. Suzuki, K. Ikegai, R. Shiraki, T. Ogiyama, T.Murakami, E. Kurosaki, A. Noda, Y. Kobayashi, M. Yokota, T. Koide, K. Kosakai,Y. Ohkura, M. Takeuchi, H. Tomiyama, M. Ohta, Bioorg. Discovery of Ipragliflozin (ASP1941): a novel C-glucoside with benzothiophene structure as a potent and selective sodium glucose co-transporter 2 (SGLT2) inhibitor for the treatment of type 2 diabetes mellitus. Med. Chem. 20 (2012) 3263–3279.

3. S.L. Schwartz, B. Akinlade, S. Klasen, D. Kowalski, W. Zhang, W. Wilpshaar. Safety, Pharmacokinetic, and Pharmacodynamic Profiles of Ipragliflozin (ASP1941), a Novel and Selective Inhibitor of Sodium-Dependent Glucose Co-Transporter 2, in Patients with Type 2 Diabetes Mellitus. Diabetes Technol. Ther. 13 (2011) 1219–1227.

4. S.A. Veltkamp, T. Kadokura, W.J. Krauwinkel, R.A. Smulders.  a novel selective sodium-dependent glucose co-transporter 2 inhibitor, on urinary glucose excretion in healthy subjects. Clin. DrugInvestig. 31 (2011) 839–851.

# 5. A. Tahara, E. Kurosaki, M. Yokono, D. Yamajuku, R. Kihara, Y. Hayashizaki, T.Takasu, M. Imamura, Q. Li, H. Tomiyama, Y. Kobayashi, A. Noda. Effects of SGLT2 selective inhibitor ipragliflozin on hyperglycemia, hyperlipidemia, hepatic steatosis, oxidative stress, inflammation, and obesity in type 2 diabetic mice. Eur. J.Pharmacol. 715 (2013) 246–255.

# 6. E. Kurosaki, H. Ogasawara. Ipragliflozin and other sodium-glucose cotransporter-2 (SGLT2) inhibitors in the treatment of type 2 diabetes: preclinical and clinical data. Pharmacol. Ther. 139 (2013) 51–59.

# 7. A. Tahara, E. Kurosaki, M. Yokono, D. Yamajuku, R. Kihara, Y. Hayashizaki, T.Takasu, M. Imamura, L. Qun, H. Tomiyama, Y. Kobayashi, A. Noda, M.Sasamata, M. Shibasaki. Pharmacological profile of ipragliflozin (ASP1941), a novel selective SGLT2 inhibitor, in vitro and in vivo. Naunyn Schmiedebergs Arch. Pharmacol. 385 (2012)423–436.

8. T. Kadokura, N. Akiyama, A. Kashiwagi, A. Utsuno, K. Kazuta, S. Yoshida, I.Nagase, R. Smulders, S. Kageyama. Pharmacokinetic and pharmacodynamic study of ipragliflozin in Japanese patients with type 2 diabetes mellitus: A randomized, double-blind, placebo-controlled study.  Diabetes Res. Clin. Pract. 106 (2014) 50–56.

9. S Kobuchi, Y Ito, K Yano, T Sakaeda. A quantitative LC–MS/MS method for determining ipragliflozin, a sodium-glucose co-transporter 2 (SGLT-2) inhibitor, and its application to a pharmacokinetic study in rats. Journal of Chromatography B1000 (2015): 22-28.

10. International conference on harmonization (ICH) of technical requirements for the registration of pharmaceuticals for human use. Validation of analytical procedures methodology. ICH-Q2 (R1), Geneva1996; 1-8.

11. ICH Stability testing of new drug substance and products Q1A (R2), International conference on Harmonization, IFPMA, Geneva 2003.

12. ICH photo Stability testing of new drug substance and products Q1B, International conference on Harmonization, IFPMA, Geneva 1996.

13. The United States Pharmacopeia, National Formulary 35, United States Pharmacopeia Convention Inc., 30th edn, 2012.

14. Quality control and quality assurance unit; Al Andalous Pharmaceutical Industries Company, Cairo, Egypt.